## Slide 1: Test Method Validation & Verification
H.P.P.S.Somasiri
Principal Research Scientist / SDD-QAD /QM
Industrial Technology Institute
SLAB
2015/11/10, BMICH
VER 6

## Slide 2: Introduction
An analytical method is the series of procedures from receipt of a sample to the production of the final report.
Validation is the process of verifying that a method is fit for purpose.

## Slide 3: Basic Concepts in Method Validation
Specificity
Analyte Stability
Analytical Range
Calibration Range
Recovery (Accuracy)
Precision
Limit of Detection (LOD)
Limit of Quantification (LOQ)
Measurement uncertainty

## Slide 4: Specificity
Shows that the detected signal is due to the analyte, not another compounds/ s.
Use matrix blanks, reagent blanks for conformation of specificity

## Slide 5: Example: Confirmation of Identity
Notes: Here is an example why two different methods are required – On the left sample and the standard perform the same – On the right when using our second technique we’ve detected a second peak and seen a reduction in peak height, an interfering compound is present.
From slide 97:
For example, the chromatography could be repeated using a column of different polarity. Where a peak is due to more than one compound, a different polarity column is a good way of separating the compounds as this column may have different selectivity. (Go to slide 101 for example chromatograms showing this)
The chromatograms on this slide illustrate this example of using different analytical columns to verify the selectivity of your test method.
When a non-polar column is used, the chromatograms just shows a single peak for both the standard and the sample and it appears as though there are no significant interferences in the sample. However, when a polar column with an alternative stationary phase is used, the standard solution still shows a single peak whereas the chromatogram corresponding to the sample shows two peaks. That is, the alternative column has been able to resolve interfering species from the target analyte, which will then lead to more accurate quantification.

## Slide 6: Analyte Stability
Analyte stability during sample storage, analyte stability in sample processing, analyte stability in sample extract should be studied.

## Slide 7: Analytical Range
Range of the analyte that can be analysed with the proposed selected method.

## Slide 8: Calibration Range
If a linear calibration is used, the range in which the response for the anayte shows a linear correlation with the analyte’s concentration should be determined.
This is known as linear dynamic range.

## Slide 9: Linear Dynamic Range

## Slide 10: Recovery
% Recover of the anlalyte on extraction, clean-up, derivatization and measurement.
Can be done with
      Certified Reference Material (CRM,s)
      Standard Reference materials (SRM,s)
      Interlaboratory comparisons
      Spiking (@ 20 %, 50 % and 80 %)

## Slide 11: Accuracy (Trueness)
How close the result to the expected (assigned) value

## Slide 12: Example: Accuracy
Sample | Blank value (mg/L) | RM value (mg/L)
1 | 0.001 | 1.46
2 | 0.018 | 1.80
3 | 0.001 | 1.17
4 | 0.018 | 1.20
5 | 0.002 | 1.22
6 | 0.002 | 0.74
7 | 0.002 | 1.16
8 | 0.014 | 1.23
9 | 0.001 | 1.20
10 | 0.002 | 1.29
Mean value | 0.004 | 1.25
RM assigned value: 1.22 ± 0.08 mg/L; Laboratory value for RM: 1.246 mg/L
Mean laboratory Value from
RM value certificate
1.246 mg/L 1.22 mg/L
Bias
value
Notes: This is an example showing the bias of a method as determined by measuring a CRM.

## Slide 13
Accuracy versus Precision
Accuracy (Trueness)
Precision
Accuracy &
Precision

## Slide 14: Precision
Is measured as “repeatability “ or reproducibility
Repeatability
Measurement is repeated with a minimum of variations (same person, same laboratory, same equipment, over a short period of time.
Reproducibility
Measurement is repeated with as much variation as possible. (Different analyst, different laboratory, different equipment, over a long period of time).

## Slide 15: Limit of Detection (LOD)
Point where the variability of the measurement is such that it is impossible to distinguish between random fluctuation in the value measured and actual signal.
  LOD = YB + 3XSB
    YB = Mean value measured from the blank
    SB = Standard Deviation of the value measured from the blank

## Slide 16: Limit of Detection
Replicate (n) | Instrument response | Calculated Concentration (ppb)
1 | 29550 | 0.100
2 | 28653 | 0.097
3 | 29401 | 0.099
4 | 28532 | 0.097
5 | 29712 | 0.101
6 | 29983 | 0.101
7 | 28167 | 0.095
Mean |  | 0.099
StdDev |  | 0.002
3 Times StdDev |  | 0.007
Detection Limit |  | 0.106
TBT in Biota
Notes: This slide shows the data…note there has been some rounding which is why the arithmetic looks a little strange.
The detection limit was calculated as the mean sample blank value plus 3 times the standard deviation of the results. That is, the mean sample blank value was 0.099 ppb and 3 times the standard deviation was 0.007 ppb. The detection limit was calculated as 0.099 ppb plus 0.007 ppb resulting in 0.106 ppb.

## Slide 17: Limit of Quantification (LOQ)
Point where the measured value could be quantified reliably
  LOQ = YB + 10XSB
    YB = Mean value measured from the blank
    SB = Standard Deviation of the value measured from the blank

## Slide 18
Example:
TBT in Biota
Replicate (n) | Instrument response | Calculated Concentration (ppb)
1 | 29550 | 0.100
2 | 28653 | 0.097
3 | 29401 | 0.099
4 | 28532 | 0.097
5 | 29712 | 0.101
6 | 29983 | 0.101
7 | 28167 | 0.095
Mean |  | 0.099
StdDev |  | 0.002
3 Times StdDev |  | 0.007
Detection Limit |  | 0.106
Limit of Quantitation
Blk + 5s 0.110
Blk + 6s 0.113
Blk + 10s 0.122
Limit of Quantification (LOQ)
Notes: The example, note that this is using the same blank data as was used for the LoD determinations.
The example on this slide which we looked at previously illustrates how to calculate the LOQ for some data obtained for tributyltin concentrations in biota samples.
In this case, a sample blank was analysed 7 times. The mean sample blank concentration was 0.099 ppb and the standard deviation of the results was 0.002 ppb.
The LOQ has been calculated as the mean sample blank value plus either 5, 6 or 10 times the standard deviation of the results to yield results of 0.110 ppb, 0.113 ppb and 0.122 ppb, respectively.

## Slide 19: Sensitivity- Cadmium Analysis by AAS
0.00
0.05
0.10
0.15
0.20
0.25
0.30
0.00
0.20
0.40
0.60
0.80
1.00
1.20
Cd Concentration (mg/L)
Absorbance
slope = 0.241 ± 0.011
intercept = 0.009 ± 0.003
r-square = 0.99442
X
X
x
x
x
x
Sensitivity
Notes: This slide shows a calibration curve constructed for cadmium analysis by atomic absorption spectroscopy.
The slope of the linear curve is 0.241 which is a measure of the instrument sensitivity.

## Slide 20: Measurement Uncertainty
Uncertainty estimates should take account of:
Overall long-term precision
Bias and its uncertainty, including the statistical uncertainty involved in the bias measurements, and the reference material or method uncertainty
Calibration uncertainties
Any significant effects operating in addition to those listed above
Notes: MU is an important part of validation, and many of the perfromance characteristics gathered as part of the validation will be used in making an estimate of MU. More discussion on MU will be covered in day 2.

## Slide 21: Test Method Verification
Any standard method shall be verified before being used in the laboratory
Verify the
Recovery
Accuracy
Precision (Repeatability and Reproducibility)
LOD and LOQ
Notes: MU is an important part of validation, and many of the perfromance characteristics gathered as part of the validation will be used in making an estimate of MU. More discussion on MU will be covered in day 2.

## Slide 22
Thank you