## Slide 1: USE OF UNCERTAINTY OF MEASUREMENT IN TESTING
Rohan Perera
MSc ( UK ), ISO/IEC 17025 Technical Assessor , Metrology Consultant

## Slide 2: Reason for estimating uncertainty
The estimation of the uncertainty of measurement allows
meaningful comparison of similar results from different laboratories or within the same laboratory
comparison of results with reference values given in specifications or standards
the equivalence of results to be judged by the user and avoid unnecessary and expensive repetition of test if differences are not significant
Uncertainty of measurement is the unifying principle in data quality

## Slide 3: Reason for estimating uncertainty
Comparison of measurements from different batches of material will not indicate real differences in properties of performance if the observed differences could simply be within the range of the inherent variation in the test procedure.
A deviation from a specified value for the properties or performance of a product may not be significant if the difference between the specified value and the test result is within the range of uncertainty.
The uncertainty of the results of a test needs to be taken into account when interpreting those results

## Slide 4: Reason for estimating uncertainty
In some cases the uncertainty of measurement or test result may be quite small or insignificant. However, without a formal estimate, one might not be able to give a convincing response if the results are challenged.
Even if the uncertainties are considered to be too small, they need to be evaluated

## Slide 5: Reason for estimating uncertainty
The results of some types of test are subjected to large uncertainty where tests are carried out on samples that are themselves inconsistent in their properties. In such cases it might be asserted that even a relatively large uncertainty to be associated with the measurement may be ignored in comparison with the uncertainty due to sample variation.
However, unless an estimate of the measurement uncertainty is made, the validity of such an assertions cannot be supported.
There can be instances where even large uncertainties may be ignored

## Slide 6: Reason for estimating uncertainty
An estimation of the components contributing to the overall uncertainty of measurement or test result provides a means of establishing that the measurements made and the results obtained are valid. It may also help to confirm the tolerances included within a performance specification have been met or that the item under test is fit for the intended purpose.
Results A & C may need further investigation and perhaps retesting
Result B can be considered as acceptable
A properly evaluated uncertainty helps to decide whether the item under test is fit for the intended purpose.
IT FACILITATES DECISION MAKING
Chart: 
| Result A | Result B | Result C
Measurement results | 10 | 10 | 10 | 10 | 10
 | 2 | 2 | 2 | 2 | 2
 | 10.5 | 5 | 2.8

## Slide 7: Reason for estimating uncertainty
A thorough assessment of the components contributing to the overall uncertainty helps to improve the procedures and accuracy of measurement or test results.
An assessment of the components contributing to the measurement uncertainty may indicate aspects of a test method to which attention should be directed in order to improve the procedure and accuracy of the measurement result and to improve the understanding of the principles of the test method. It plays a key part in method validation.
..\LCS\Test Methods\Calibration of Balance LCS_TM_09\LCS_WS_09_Rev0.xlsx
Source of uncertainty | Variability interval | Standard uncertainty | Sensitivity coefficient | Contribution to uncertainty
Balance resolution | 0.0028 mg | 0.000 40 mg | 1 | 0.000 40 mg
Balance repeatability | 0.008 mg | 0.002 31mg | 1 | 0.002 31 mg
Calibration of standards | 0.0082 mg | 0.002 05 mg | 1 | 0.002 05 mg
Drift of standards | 0.02 mg | 0.005 77 mg | 1 | 0.005 77 mg
Temperature fluctuation | 1 ⁰C | 0.289 ⁰C | 0.01 mg/⁰C | 0.002 89 mg
Off-centre loading | 5 mm | 1.44 mm | 0.000 4 mg/mm | 0.000 58 mg
Expanded uncertainty (k=2) |  |  |  | 0.016 mg

## Slide 8: Reason for estimating uncertainty
There is a general requirement that all accredited laboratories shall have a procedure in place to evaluate the measurement uncertainty of measurement or test results for which the laboratory holds accreditation. The degree of rigour needed in an estimation of uncertainty may differ between testing and calibration laboratories. In some cases the nature of the test method may preclude rigorous and metrologically and statistically valid calculation of measurement uncertainty. However, the laboratory is required to attempt to identify all possible sources of uncertainty and make a reasonable estimate.
THANK YOU
Iso/iec 17025
THE STANDARD REQUIRES THAT ALL ACCREDITED LABORATORIES “shall have and shall apply a procedure to estimate the uncertainty of measurement”